## Inputs
Inputs
Editable constants used by the calculation tables.
λ_nm | Label | PDC_mW | ModDepth | ΔPopt_mW | μeff_cm^-1 | RD_A/W | ZT_Ω | RL_Ω | Pnoise_2cm_dBm | Pnoise_4cm_dBm
670 | 670 nm | 1.5 | 0.2 | 0.3 | 2.31 | 0.65 | 794 | 50 | -135.95 | -143.63
780 | 780/800 nm | 0.5 | 0.2 | 0.1 | 2.04 | 0.65 | 794 | 50 | -138 | -143.64
850 | 850 nm | 2 | 0.5 | 1 | 2.17 | 0.65 | 794 | 50 | -133.56 | -143.38
Frequency_MHz | Frequency_Hz | ω_rad_s |  | Depth_cm
300 | 300000000 | 1884955592.1539 |  | 2
400 | 400000000 | 2513274122.8718 |  | 4
500 | 500000000 | 3141592653.5898
600 | 600000000 | 3769911184.3078
I/Q correction | Value
Conversion_loss_dB | 6
Noise_figure_dB | 6
Total_IQ_SNR_correction_dB | 12

## Part 2 Frequency
Part 2. Modulation Frequency
Formula: ω = 2πf. This table converts frequency to angular frequency.
Frequency | Frequency_Hz | ω_rad_s
300 MHz | 300000000 | 1884955592.1539
400 MHz | 400000000 | 2513274122.8718
500 MHz | 500000000 | 3141592653.5898
600 MHz | 600000000 | 3769911184.3078

## Part 3 Tx Power
Part 3. Transmitter Optical Power
Formulas: ΔPopt = mPDC and Ptx(t,f)=PDC+ΔPopt cos(2πft).
Wavelength | PDC_mW | ModDepth | ΔPopt_mW | Ptx waveform
670 nm | 1.5 | 20% | 0.3 | 1.5+0.30cos(2πft) mW
800 nm | 0.5 | 20% | 0.1 | 0.5+0.10cos(2πft) mW
850 nm | 2 | 50% | 1 | 2.0+1.00cos(2πft) mW

## Part 4 dBm
Part 4. Optical Power in dBm
Formula: PdBm = 10log10(PmW).
Wavelength | PDC_mW | PDC_dBm
670 nm | 1.5 | 1.7609
800 nm | 0.5 | -3.0103
850 nm | 2 | 3.0103

## Part 5 Elec Efficiency
Part 5. Laser Electrical Power and Efficiency
Formulas: Pelec=VFIF and η=Poptical/Pelec×100%.
Wavelength | VF_V | IF_mA | Pelec_mW | Poptical_mW | Efficiency_%
670 nm | 2.1 | 4 | 8.4 | 1.5 | 17.8571
800 nm | 1.9 | 5 | 9.5 | 0.5 | 5.2632
850 nm | 1.8 | 6.5 | 11.7 | 2 | 17.094

## Part 6 Linearity
Part 6. Laser Current Linearity Check
Checks that laser current stays above threshold and below maximum current.
Wavelength | Min current_mA | Max current_mA | Ith_mA | Imax_mA | Result
670 nm | 4.5 | 6 | 3.5 | 7 | Pass
800 nm | 4.5 | 5.5 | 4 | 8 | Pass
850 nm | 4 | 9 | 1.5 | 12 | Pass

## Part 7 Attenuation
Part 7. Tissue Effective Attenuation
Formula: μeff = sqrt(3μa(μa+μ's)).
Wavelength | μeff_cm^-1
670 nm | 2.31
780/800 nm | 2.04
850 nm | 2.17

## Part 8 Tissue Loss
Part 8. Tissue Loss for 2 cm and 4 cm
Formula: Ltissue,dB=4.343μeffd. Chart compares shallow and deeper path loss.
Wavelength | μeff_cm^-1 | Loss_2cm_dB | Loss_4cm_dB | Extra_loss_dB
670 nm | 2.31 | 20.0647 | 40.1293 | 20.0647
780/800 nm | 2.04 | 17.7194 | 35.4389 | 17.7194
850 nm | 2.17 | 18.8486 | 37.6972 | 18.8486

## Part 9 Detected Power
Part 9. Detected Optical Power
Formulas: Pdet,DC=PDCe^(-μeffd), ΔPdet=ΔPopte^(-μeffd).
Wavelength | Pdet_DC_2cm_uW | ΔPdet_2cm_uW | Pdet_DC_4cm_uW | ΔPdet_4cm_uW
670 nm | 14.7792 | 2.9558 | 0.14562 | 0.029123
800 nm | 8.4537 | 1.6907 | 0.14293 | 0.028586
850 nm | 26.0731 | 13.0365 | 0.3399 | 0.16995

## Part 10 Receiver
Part 10. Receiver Current, Voltage, and Signal Power
Formulas: I=RDPdet, V=IZT, Psig=VAC²/(2RL).
Wavelength | Depth_cm | Iavg_uA | Vout_DC_mV | IAC_uA | VAC_mV | Psig_dBm
670 nm | 2 | 9.6065 | 7.6275 | 1.9213 | 1.5255 | -46.3317
670 nm | 4 | 0.094651 | 0.075153 | 0.01893 | 0.015031 | -86.4605
780/800 nm | 2 | 5.4949 | 4.363 | 1.099 | 0.87259 | -51.1838
780/800 nm | 4 | 0.092905 | 0.073767 | 0.018581 | 0.014753 | -86.6222
850 nm | 2 | 16.9475 | 13.4563 | 8.4737 | 6.7282 | -33.4421
850 nm | 4 | 0.22094 | 0.17542 | 0.11047 | 0.087712 | -71.1388
Wavelength | Psig 2 cm (dBm) | Psig 4 cm (dBm)
670 nm | -46.3317 | -86.4605
780/800 nm | -51.1838 | -86.6222
850 nm | -33.4421 | -71.1388

## Part 11 SNR
Part 11. Receiver Noise and SNR
Formula: SNRdB=Psig,dBm−Pnoise,dBm. Post-IQ subtracts 12 dB.
Wavelength | Depth_cm | Pnoise_dBm | Psig_dBm | SNR_preIQ_dB | SNR_postIQ_dB
670 nm | 2 | -135.95 | -46.3317 | 89.6183 | 77.6183
670 nm | 4 | -143.63 | -86.4605 | 57.1695 | 45.1695
780/800 nm | 2 | -138 | -51.1838 | 86.8162 | 74.8162
780/800 nm | 4 | -143.64 | -86.6222 | 57.0178 | 45.0178
850 nm | 2 | -133.56 | -33.4421 | 100.1179 | 88.1179
850 nm | 4 | -143.38 | -71.1388 | 72.2412 | 60.2412

## Part 12 IQ Phase
Part 12. I/Q Demodulation and Phase
Formulas: A=sqrt(I²+Q²), φ=tan⁻¹(Q/I). Phase trend estimated using τ=nd/c and φ=2πfτ.
Depth_cm | Frequency_MHz | Delay_s | Phase_rad | Phase_deg_unwrapped | Phase_deg_mod360
2 | 300 | 9.3333e-11 | 0.17593 | 10.08 | 10.08
2 | 400 | 9.3333e-11 | 0.23457 | 13.44 | 13.44
2 | 500 | 9.3333e-11 | 0.29322 | 16.8 | 16.8
2 | 600 | 9.3333e-11 | 0.35186 | 20.16 | 20.16
4 | 300 | 1.8667e-10 | 0.35186 | 20.16 | 20.16
4 | 400 | 1.8667e-10 | 0.46914 | 26.88 | 26.88
4 | 500 | 1.8667e-10 | 0.58643 | 33.6 | 33.6
4 | 600 | 1.8667e-10 | 0.70372 | 40.32 | 40.32

## Part 13 MBLL
Part 13. Optical Density and MBLL
Formula: ΔOD=log10(I0/Iλ), then MBLL relates ΔOD to Δ[HbO2] and Δ[Hb].
Wavelength | εHbO2 | εHb
670 nm | 0.000294 | 0.002795
780 nm | 0.00071 | 0.001075
850 nm | 0.001058 | 0.000691
800 nm approximation | 0.000816 | 0.000762

## Part 14 Leakage
Part 14. Leakage Location Interpretation
Higher absorption increases μeff, increases tissue loss, and lowers detected intensity.
Interpretation step | Effect
Blood leakage increases absorption | ↑ μa
Higher absorption increases attenuation | ↑ μeff
Higher attenuation increases loss | ↑ Ltissue
Higher loss lowers detected intensity | ↓ Iλ / ↓ Pdet
2 cm path | Stronger shallow signal
4 cm path | Weaker deeper signal

## Dashboard
Dashboard Summary
Quick comparison from charted calculation sections.
Wavelength | Loss_2cm_dB | Loss_4cm_dB | Pdet_DC_2cm_uW | Pdet_DC_4cm_uW | PostIQ_SNR_2cm_dB | PostIQ_SNR_4cm_dB
670 nm | 20.0647 | 40.1293 | 14.7792 | 0.14562 | 77.6183 | 45.1695
780/800 nm | 17.7194 | 35.4389 | 8.4537 | 0.14293 | 74.8162 | 45.0178
850 nm | 18.8486 | 37.6972 | 26.0731 | 0.3399 | 88.1179 | 60.2412